| Company Name/Logo | Water Activity Calibration Verification Record Example | Version No.: 01 Last revised on: mm-dd-yyyy Approved by: [Personnel responsible] Effective Date: mm-dd-yyyy |
| --- | --- | --- |

| Date | Verification Standard | Expected Water Activity | Observed Water Activity | Retest (N/A if not required) | Initials | Comments |
| --- | --- | --- | --- | --- | --- | --- |
|  | 6.00 mol/kg of NaCl | 0.72± 0.02* |  |  |  |  |
|  | 2.33 mol/kg of NaCl | 0.92± 0.02 |  |  |  |  |
|  | 6.00 mol/kg of NaCl | 0.72± 0.02 |  |  |  |  |
|  | 2.33 mol/kg of NaCl | 0.92± 0.02 |  |  |  |  |
|  | 6.00 mol/kg of NaCl | 0.72± 0.02 |  |  |  |  |
|  | 2.33 mol/kg of NaCl | 0.92± 0.02 |  |  |  |  |

If the Aw meter reads outside this range, the sample chamber will need to be recleaned following manufacturer’s instructions and reverified. If this does not work inform Operator and the unit may require service. Verification will be done once per shift or daily before use. See calibration procedure for further details. * The accuracy value should be according to the manufacturer’s instructions.

Verification date and initials: ________________ Result:      acceptable        unacceptable-requires follow-up

Was this an on-site verification?      Yes      No      Result:      acceptable        unacceptable-requires follow-up